## List of Items
EQUIPMENT REQUIRED -32 CHANNEL DRY ELECTRODE
S.No | List of Equipment | Description | Specifications | Qty Required | Inclusive items | Price | Total
[6 empty] | (INR) | (INR)
1 | 32 Channel dry electrode wireless EEG headset | 32 channel Dry Electrode wireless EEG Headset with integrated EEG amplifier and wireless electronics inbuilt in the EEG headset which includes 30 fixed channels —10-20 montage + 10 additional on-head channels — plus 2 ExG channels for gathering additional biometric data | ● 32 channel light weight, dry electrode wireless EEG amplifier headset for fast and zero-gel recordings headsets. ● EEG amplifier and wireless electronics should be integrated into the headset. ● built-in noise rejection and LED impedance indicators for real-time contact quality. ● flexible composite arms for easy movement & excellent contact between sensors and scalp. ● 2 variable ExG leads for user-defined EEG ECG, EMG, and EOG. ● 8 hours of uninterrupted data recording with two AA batteries. ● Features: ○ A/D Resolution: 24-bit ○ Sampling rate: 500Hz ○ Bandwidth: 0-131 Hz with DC coupling. ○ 3-axis accelerometer ○ Wireless Range: 10 meters. ○ Noise: <1.0 µV RMS from 1-50 Hz, shorted inputs. ○ Weight: less than 650g ○ Fits heads sized 52-62 cm ● Data streaming via Bluetooth Low Energy. Should have full access to raw data via real-time streaming API. ● Continuous on-board and on-screen impedance check with real-time monitoring of all channels simultaneous with EEG. ● compatible with LabStreamingLayer, EEGLAB, BCILAB, MATLAB, BCI2000, OpenViBE. ● offered with Windows based Recording Software, should have wizard based workspaces for various research requirements and protocols. Menu structure to guide the user with all crucial steps for setting up the hardware/software filter configuration, selecting the hardware filters on a channel by channel basis. ● Data Recording Software able to record and process the data acquired from electrodes and auxiliary sensors with the below features. ○ compatible with amplifier to ensure maximum compatibility and a better sampling rate. Open source or universal processing software like Matlab or toolboxes will not be accepted. ○ The selection of hardware filters on a channel-by-channel basis should be easy and fast. The acquired data should be displayed in multiple ways and the channel montages (original, bipolar and average) to be made switchable on the fly, to adjust the channel view to the specific experiment needs. ○ check the impedance value channel by channel and display with a fully selectable colour coding. ○ perform a complete evoked potential analysis in real-time directly and the segmented/averaged data should be stored together with the raw data. ○ compatible to send out the incoming data to the network via the TCP/IP protocol using the Remote Data Access module for real-time data analyses. ○ Online Average Overlay-allows you to add a static overlay of a previously recorded average to the average view and thus compare the ongoing wave with a prototypical one. ○ selection of Default Display Interval i.e to select the length of the time interval to be displayed by default on the computer screen.○ select a subset of channels during monitoring. ○ Interface to the active electrode control software, should be able to automatically save the impedance values in the header file of EEG data and should not need a separate file to save them. ○ transfer the data to the other Programs like MATLAB and Python, located on the local computer or on computers in the network. ○ BCI Applications provision for scripting and developing user defined analysis for BCI tools along with Data exporting. ○ The data scripting is enabled where the user can script their own code for special purpose applications along with the export facility to load into any other analysis tools is available. | 1 | 30 Drypad Sensors 60 Flex Sensors 10 Drypad Ear Sensors 30 Skintact Sensors A1 Earclip A2 Earclip 3 Active Lead Wires 2 Passive Lead Wires 5 Alcohol Wipes 4 Rechargeable AA Batteries Battery Charger Carrying Case 3 Year Warranty
Grand Total (INR)